Customer Product Quality Requirements Form

Part Number: OXY4823		    X      New Product       	               Revised Product

Company: OXY Whichita		Contact:  Dale Jones

Address:

Phone:					email:  Dale_Jones@oxy.com

*If not already indicated, please specify if this product needs to be prepared or verified to an

industry standard procedure as ASTM, EPA, APHA or other method:

Note: in the event that the above method becomes revised, it will be the customer’s responsibility to provide updated information / revision requests to Aqua Solutions.

OXY4823       	STARCH INDICATOR,  SPECIAL FOR ANALYSIS OF SODIUM HYPOCHLORITE

RINSE VESSEL WITH DEIONIZED WATER  BEFORE USE.  INSPECT STIR BAR FOR METAL AND REMOVE IF NECESSARY!!! NOTE THIS IS A CUSTOMER SUPPLIED METHOD AND CANNOT BE ALTERED WITHOUT THE PERMISSION OF OCCIDENTAL!

PREPARE THIS SOLUTION VOLUMETRICALLY.

NOTE: DO NOT MAKE MORE THAN 6 LITERS AT A TIME.

1 LITER FORMULATION

FOR EACH LITER OF FINALSOLUTION, MAKE A PASTE USING 6.0 GRAMS OF SOLUBLE POTATO

STARCH (S5222, STARCH REAGENT ACS) AND A SMALL AMOUNT OF DEIONIZED WATER (9765). ADD

THE PASTE TO 990 MLS OF  BOILING DEIONIZED WATER AND STIR UNTIL HOMOGENOUS.  ONCE

COOLED, ADD 20 MLS OF POTASSIUM HYDROXIDE 45% (7285) AND 28 MLS OF HYDROCHLORIC ACID,

REAGENT ACS GRADE (H2505).  ADJUST THE PH TO 6.0 +- 0.2 WITH KOH OR HCL.  RECORD THE PH

HERE _______________________________AFTER ADJUSTING THE PH, ADD 6 MLS OF REAGENT ACS

GRADE GLACIAL ACETIC ACID (A1100) AS A PRESERVATIVE.  STIR UNTIL HOMOGENOUS. AFTER

ADDING  THE ACETIC ACID, AND BEFORE RELEASE, CHECK THE FINAL PH OF THE SOLUTION PREVIOUS

BATCHES HAVE BEEN AROUND 2.9. RECORD THE PH AT TIME OF RELEASE

HERE:________________________ SHOULD BE AROUND 2.9 +- 0.5

VERIFICATION

ADD 2-4 MLS OF PRODUCT TO 10 MLS OF 0.1 NORMAL IODINE (4750). THE COLOR SHOULD BE REDDISH

BLACK. ADD SODIUM THIOSULFATE, 0.1 NORMAL(S3999) DROPWISE UNTIL THE SOLUTION CLEARS.

SHOULD TAKE 10.0 +- 0.1 MLS.

REVISION:  0.6	    DATE:  06/26/2019                     INITIAL:STN

EXPIRATION: 3 MONTHS

SPECIAL LABELING INSTRUCTIONS: NONE

FILTERING INSTRUCTIONS: NONE

SPECIAL PACKAGING INSTRUCTIONS: DO NOT PUT STARCH DOWN THE LAB SINKS

PACKAGING: OXY4823 = 1 LITER AMBER PLASTIC WITH FLIP TOP CAP, OXY4823A = 250 ML AMBER

PLASTIIC WITH FLIP TOP CAP

HMIS CODE:  2OOC

PERSONAL PROTECTION: Nitrile GLOVES, SPLASH GOGGLES, LABJACKET

STANDARDIZATION: INITIALLY ADJUST THE Ph TO 6.000 +- 0.200 WITH 45% POTASSIUM HYDROXIDE THEN AFTER PH IS STABLE, ADD 6 MLS OF ACETIC ACID PER LITER. FINAL Ph SHOULD BE 2.900 +- 0.500  USE ABOVE VERIFICATION PROCEDURE TO VERIFY THE INDICATOR WILLWORK PROPERLY! Finally, use densitometer to verify density of 1.01618± 0.00198@20°C (1.01420- 1.01817)

APPEARANCE OF FINAL SOLN/ ODOR: CLEAR TO SLIGHT HAZY  /  ACETIC ACID SMELL

DISPOSAL: GENERAL

REVISION 0.1, 07/07/05: CHANGED PART NUMBER TO S5222 FROM S5261, BOTH ARE POTATO STARCH. STN

REVISION 0.2, 11/09/05: CHANGED EXPIRATION DATE TO 3 MONTHS FROM 1 YEAR AFTER PROBLEM OCCURRED. STN

REVISION 0.3, 10-31-13: REVISED INSTRUCTIONS TO RECORD FINAL PH TO INSURE THAT THE ACETIC

ACID IS ADDED AFTER THE PH ADJUSTMENT TO 6.0. THIS WILL MAKE PH LOWER THAN 6.0 AND ABOUT

2.9. STN

REVISION 0.4: 09/08/14,  ADDED INSTRUCTIONS TO  PREPARE VOLUMETRICALLY & TO NOT MAKE MORE

THAN 6 LITERS AT A TIME. UPDATED TO NOTE9.. STN

Revision 0.5, 10-16-2015: added density information. STN

Revision 0.6,06/26/19: updated density information based on at least 8 data points. STN

***Please make any revisions, sign and email or fax back to Aqua Solutions @ 281-479-1833

Customer Authorization Signature:  ________________________________________  Date:         /           /

Laboratory Approval: __________________________________________________  Date:       /         /