Customer Product Quality Requirements Form

Part Number: 1440		    X      New Product       	               Revised Product

Company: OXY Whichita		Contact:  Dale Jones

Address:

Phone:					email:  Dale_Jones@oxy.com

*If not already indicated, please specify if this product needs to be prepared or verified to an

industry standard procedure as ASTM, EPA, APHA or other method:

Note: in the event that the above method becomes revised, it will be the customer’s responsibility to provide updated information / revision requests to Aqua Solutions.

1440			BUFFER SOLUTION pH 2.00 +- 0.010 @25° C

N.I.S.T. TRACEABLE

SPECIAL PREPARATION INSTRUCTIONS: CLEAN THE GLASSWEAR WITH DIH20 (9765) BEFORE STARTING THE BATCH. WEAR GOGGLES AND GLOVES WHEN MAKING THE SOLUTION.  NOTE: THIS PRODUCT IS USED AS A REFERENCE BUFFER AND IT IS CRITICAL THE THE PH BE 2.00 +- 0.010 @25 DEG C.

1 LITER FORMULATION

FOR EACH LITER OF FINAL SOLUTION, DISSOLVE 3.728 GRAMS OF REAGENT ACS GRADE POTASSIUM CHLORIDE (P3550) INTO ABOUT 800 MLS OF DEIONIZED WATER (9765). ADJUST THE PH TO 2.00 +- 0.010 @ 25° WITH 50% V/V HYDROCHLORIC ACID (4420).  IT IS CRITICAL THAT YOU BRING THE TEMPERATURE TO 25 DEG C +- 0.5 BEFORE ADJUSTING THE PH.  DILUTE TO 1 LITER WITH DEIONIZED WATER AND STIR UNTIL HOMOGENOUS.CHECK AND IF NECESSARY, READJUST THE PH TO 2.00 +- 0.010 @ 25 DEG C. WITH 0.1 NORMAL HCL (4130) OR 0.1 NORMAL SODIUM HYDROXIDE (8625). Note: this product is certified using QCSP20. you must prepare NIST buffer 1.68 (srm 189) and 4.01 buffer (srm 185)  to calibrate the meter with. Specifications are 2.00 +- 0.01 @25 deg C.The buffer must be checked and within specifications for 2 consecutive days before able to be released.   When 	updating the pH meter, update using the 4.00 Buffer. Use NIST preparation form and the attached pH meter calibration form for reference buffers when preparing this product. keep original with paperwork and put a copy in the appropriate pH meter log book.  Once in specifications, calibrate the pH meter with the new lot of 2 buffer and  pH 4 buffer. To be acceptable for release, the slope must be between 98.5 – 101%

REVISION:  2.5          DATE:  08-06-2014           INITIALS:  STN

EXPIRATION: 1 YEAR

SPECIAL PACKAGING: NONE

PACKAGING: 10 LITER CUBE; 1 LITER CLEAR PLASTIC; 20 LITER CUBE;

4 LITER CLEAR PLASTIC; 4 LITER CUBE; 500 ML CLEAR PLASTIC

HMIS CODE: 100C

PERSONAL PROTECTION: GLOVES AND GOGGLES

STANDARDIZATION: QCSP20 to pH 2.000 +/- 0.010 @ 20 Deg C The buffer must be checked and within specifications for 2 consecutive days before able to be released.   When updating the pH meter, update using the 4.00 Buffer. Use NIST preparation form and the attached pH meter calibration form for reference buffers when preparing this product. keep original with paperwork and put a copy in the appropriate pH meter log book.  Once in specifications, calibrate the pH meter with the new lot of 2 buffer and pH 4 buffer. To be acceptable for release, the slope must be between 	 98.5 – 101%.THE SLOPE MUST BE 98.5 - 101%. RECORD SLOPE

HERE:__________________ INITIAL:_______________

APPEARANCE/ODOR: CLEAR WATER WHITE LIQUID / NONE  INITIALS:  ________

DISPOSAL: GENERAL WASTE

REVISION: 2.0, 12/06/99, REVIEWED AND REVISED BATCH SHEET AND BATCH CARDTO SAME FORMULATION. STN

REVISION: 2.1, 09/23/08, INCLUDED INSTRUCTION TO CHECK pH AFTER DILUTING

TO FINAL VOLUME AND, IF NECESSARY, ADJUSTING TO 2.00 WITH SODIUM HYDROXIDE OR HYDROCHLORIC ACID AFTER A PROBLEM OCCURRED. STN

REVISION: 2.2, 08/03/09, ADDED 0.5 GRAMS OF DOWICIL TM 75 D7305 TO THE

SOLUTION AND CHANGED SPECIFIACATIONS ON THE pH FROM 2.00 +/- 0.02 TO

2.000 +/- 0.010. STN

REVISION: 2.3, 08/24/09, REMOVED THE DOWICIL FROM THE SOLUTION BECAUSE IT WAS TURNING IT YELLOW. STN

REVISION 2.4: INCLUDED SPECIFIC INSTRUCTIONS TO CALIBRATE PH METER WITH NIST 1.68 AND 4 AND TO VERIFY A GOOD SLOPE WITH NEW 2 AND 4 ONCE IN SPEC. STN

Revision 2.5, included specific use of Calibration form for reference buffers which will print out with batch sheet. STN

***Please make any revisions, sign and email or fax back to Aqua Solutions @ 281-479-1833

Customer Authorization Signature:  ________________________________________  Date:         /           /

Laboratory Approval: __________________________________________________  Date:       /         /